Voltammetric determination of Sulfaclozine Sodium at Sephadex-modified carbon paste electrode

Emad M. Hussiena*, Hanaa Salehb, Magda El Henaweeb, Afaf Abou El Khairband Neven Ahmeda

aNational Organization of Drug Control and Research (NODCAR), Giza, Egypt.

bFaculty of Pharmacy, Zagazig University, Zagazig, Egypt.

*Correspondence author at: National Organization for Drug Control & Research, P.O Box 29, Giza, Egypt. Tel.: +2 02 3749 6077.

E-mail address: emadhussien@yahoo.com (Emad M. Hussien)

Abstract

The electrochemical oxidation of Sulfaclozine Sodium (SCL) was studied at a bare and sephadex-modified carbon paste electrodes by cyclic voltammetry and square wave voltammetry. Cyclic voltammetry giving a well-defined irreversible oxidation peak at 0.94 V using Britton- Robinson buffer pH 7.0. The strong affinity of Sulfaclozine Sodium to sephadex allowed accumulation of Sulfaclozine Sodium at the electrode surface and thus higher electrochemical sensitivity to Sulfaclozine Sodium. The influence of sephadex content, pH of the solution and the scan rate on the modified electrode sensitivity was studied. A linear calibration curve covering the concentration range from 5 µM to 1 mM was obtained using square wave voltammetry. The voltammetric study was successfully applied for determination of SCL in veterinary pharmaceutical formulations with satisfactory accuracy and precision.

Keywords: Sulfaclozine Sodium; square wave voltammetry; sephadex carbon paste

electrode.

# Introduction

Coccidiosis is a parasitic disease attacks the intestinal tract of poultry caused by protozoan parasites of the genus Eimeria. This disease is of worldwide occurrence and costs the poultry industry many millions of dollars every year to control.1Although live vaccines were introduced, in most countries prophylactic chemotherapy is still the preferred method for coccidiosis control. Last half of the twentieth century marked improvements in the performance of commercially reared poultry have been made. These improvements would not be possible without the introduction of a succession of ever more effective anticoccidial agents to control coccidiosis.2

Sulfaclozine, N1 - (6-chloropyrazinyl) sulfanilamide (Figure 1) is a sulfonamide antibacterial that has been used in veterinary medicine.3 it is effective in the treatment of clinical coccidiosis as well as prevention of the disease.4

Analytical methods reported for sulfaclozine determination have been chromatographic methods with different detectors.5,6,7,8,9 and capillary electrophoresis.10These methods are either time-consuming or use expensive instrumentation.

Voltammetric techniques have been found to be simple with high sensitivity and selectivity for drug analysis. Indeed, Nemours pharmaceutical compounds were analyzed with voltammetry.11The proposed method is rapid with better sensitivity compared to the well-known sophisticated expensive techniques such as LC/MS/MS. Carbon paste electrodes have several advantages including simple preparation, easy surface regeneration and modification.11,12 It has been reported that surfactants can influence the electrochemical process occurring at the surface of the electrode.13,14 Sephadex is a cross-linked dextran which is used as a stationary phase in gel filtration chromatography.15the usage of Sephadex in carbon paste electrode modification for nifuroxazide and glibenclaide determination has been reported.16,17

The present work report for the first time a SWV method for SLC determination in veterinary formulation. The proposed method utilizes the electrochemical oxidation of SLC at the carbon paste electrode surface in a micellar medium. The effect of sephadex on the oxidation peak current was investigated. The method was validated According to the ICH guidelines.18

# Experimental

2.1. Reagents and materials

Sulfaclozine sodium was obtained from Yangzhou Tianhe Pharmaceutical Co., LTD, China with potency 99.75%. Clozicocc® W.S.P(batch No.564) (Each 100 gm contains 32 gmsulfaclozine sodium) was obtained from Pharco Pharmaceuticals, Alex., Egypt. Graphite powder, paraffin oil, Sephadex G-50, C18 silica gel and chitosan were supplied from Sigma–Aldrich.Methanol was purchased from LobaChemie Co., India. Sodium dodecyl sulfate (SDS), Phosphoric acid and boric acid were supplied from Adwic Co., Egypt. Acetic acid was obtained from Piochem Co., Egypt. Briton Robinson buffer (B-R) buffer was prepared by adding equal volumes of phosphoric acid 0.04 M, acetic acid 0.04 M and boric acid 0.04 M, the pH of the buffer adjusted by NaOH 0.2M to cover the pH range from 2.0 to 10.0. SDS 10.0 mM was prepared by dissolving appropriate amount in water. Double- distilled water was used throughout the study and referred to by "water”.

2.1.4. Standard solution

Sulfaclozine stock solution (10.0mM) prepared by dissolving 30.66 mg of SLC in 1.0 mL methanol, then diluting with water to 10 mL.

2.2. Apparatus

Bio-logic SP 150 electrochemical workstation with a three-electrode configured stand (model C-3). The working electrode was a bare carbon paste electrode or a Sephadex-modified carbon paste electrode (SMCPE), Ag/AgCl (3 MKCl(BAS, USA) was the reference electrode and Platinum wire as counter electrode (BAS, USA).

2.3. Procedures

2.3.1. Preparation of modified carbon- paste electrode

Sephadex-modified carbon paste (SMCPE) electrode was made by hand mixing of 0.4g Sephadex and 0.8 g of graphite powder then 0.4 mL paraffin oil added in a mortar. Plain (unmodified) carbon paste was made by mixing 1 g graphite powder and 0.6mL paraffin oil. The prepared paste packed into the electrode body then smoothed on a filter paper till shiny appearance is obtained.

2.3.2. Analytical procedure

The CV of the sephadex-modified carbon paste was repeated between 0 and 1.4 V several times in the buffer solution (pH 7.0) till the cyclic voltammograms becomes stable. Then the electrode was transferred into another cell of buffer solution (pH 7.0) containing 0.005mM to 1.0 mM SLC and 0.03mM SDS. The electrode remains 30 s under stirring at open circuit potential, afterwards cyclic voltammogram was recorded between +0.4 and +1.4, at 100 mVs−1 scan rate.

2.3.3. Calibration curve of SLC

The square wave voltammetry (SWV) was performed to determine SLC in bulk powder and pharmaceutical formulations. Different aliquots equivalent to (0.005mM to 1.0mM) were accurately transferred from the stock standard solution to the electrolytic cell containing 10 mL buffer (pH 7.0)and 0.03mM SDS. The square wave voltammograms were recorded at SMCPE. The peak current was plotted against drug concentration of in (µM).

2.3.4. Application to veterinary pharmaceutical formulation

0.96g Clozicocc® W.S.P. (containing 306.7 mg of sulfaclozine sodium) accurately weighed and transferred to a 100 mL volumetric flask, 10 mL methanol added and sonicated for 15 min, then the volume was completed with water to obtain 10.0mM(solution 1), then 1 mL was transferred from (solution 1) into 10mL volumetric flaskand complete to volume with water to obtain 1.0mM(solution 2).

An aliquots[50µL from solution 1, 500 µL and 700 µL from solution 2] were separately transferred to the electrolytic cell containing 10 mL buffer solution of the optimum pH (7.0) and 0.03mM SDS, to obtain a final concentration [0.005, 0.47 and 0.65mM of test solution] , the drug was analyzed by the proposed voltammetric procedure using the standard addition method.

3. Results and discussion

3.1 Sulfaclozine electrochemical oxidation behavior

We use cyclic voltammetry to study the electrochemical oxidation of SLC at carbon paste electrode. The Sulfaclozine CV voltammogram (Figure 2) showing one anodic peak current ((I) = 3.5 µA at 0.94 V), and it was no cathodic peak observed in the reverse scan, it means that the oxidation process of SLC is irreversible.This oxidation current could be due to the oxidationof SLCamine group.19

3.2. Experimental conditions optimization

3.2.1. Effect of pH

Sulfaclozine potential-pH behavior was studied in BR buffers over a pH range of 2.0– 10.0. Peak potential is shifted toward less positive values when the pH was increased, The relationship between EP and pH at sephadex - modified carbon paste electrode was found to be linear (Figure 3a) and controlled by the equation EP=-51pH + 1284

(R2 = 0.995). The slope (~ 51 mV per pH) is near to the expected 59 mV per pH indicating that equal number of protons and electrons involved in the oxidation process of SLC. At pH 7.0 the highest current value and sharper response was found (Figure3b) which was chosen for subsequent measurements.

3.2.2. Effect of Sephadex

Different materials including C18 modified silica, chitosan and sephadex were tested for possible enhancement of oxidation current and, hence, increasing sensitivity of the electrode. (Figure 4) shows no difference in the electrochemical behavior of sulfaclozine when 30%C18 modified silica was added to the carbon paste electrode, while the addition of 30% chitosan to the paste make little improvement in the current response. In contrast, 30%Sephadex-modified electrode exhibited a considerable larger oxidation signal that indicating the affinity of the drug to Sephadex. This affinity can be can be utilized for drug preconcentration onto the electrode surface.

The effect of Sephadex on the peak current was investigated using different amount of sephadex in the paste (Figure 5). The highest peak current was obtained with 25% w/w of Sephadex.

3.2.3. Effect of Sodium dodecyl sulfate

Sulfaclozine oxidation behavior in a micellar medium was also studied using SDS.

SDS is a hydrophobic ionic surfactant, which can be adsorbed onto the electrode surface. As a result, the electrochemical process such as the mass and electron transfer energy at the electrode/solution interface are affected.20It has been reported that SDS can remove the oily binder (insulator) and hence lower the uncompensated resistance at the electrode/solution interface.21, 22, 23 A different volume of 0.01 M SDS (10–50 µL) was added to the electrochemical cell and the CV was recorded. The anodic current againstSDS concentrations was plotted (Figure 6). It was observed that upon SDS addition, the peak current increases proportionally, giving the highest current signal with 30 µL SDS equivalent to 0.03mM, above this concentration no further improvement in the peak current was observed.

3.2.4. Effect of scan rate

The scan rate (υ) effect on peak potential (Ep) and peak current (ip) was studied between 10mVs–1and 250mVs–1in 1.0mM sulfaclozine solution in BR buffer (pH 7.0) containing 0.03mM SDS (Figure7a),The relationship between the oxidation peak current of SLC and the square root of scan rate (υ 1/2) was found to be linear, indicating that electrochemical oxidation of SLC is a diffusion controlled process.24

Plotting log the anodic peak current against log the scan rate gave a straight line with a slope of 0.47as shown in (Figure7b), which is nearly equal to the theoretical value of 0.5 for a purely diffusion-controlled process.24It was also found that the Ep (oxidation peak potential) was dependent on scan rate, peak potential was shifted to more positive values when the scan rate increases, which confirms that the oxidation process was irreversible. (Figure 7C) shows a linear relationship between the peak potential and log the scan rate. According to Laviron׳s equation.25

Ep = E0′ + (2.303RTαnF)log ( RTk0αnF) + (2.303RTαnF)log υ                                  (1)

Where α is the transfer coefficient, k0 is the standard heterogeneous rate constant of the reaction, υ is the scan rate, and E0 is the formal redox potential, n is the number of electrons transferred. So the value of αn can be obtained from the slope of Ep vs log υ. The slope was found to be 0.068, when T = 298K and R =8.314 JK−1 mol−1 and F=96485 C/mol, αn will be equal to 0.85, According to Bard and Faulkner.26α can be given as

α = 47.7Ep-Ep/2mV                                                                                   (2)

k0 value can be calculated from the intercept of the above plot if the value of E0′ is known.

E0′ in Eq. (1) can be obtained from the intercept of Ep versus υ curve by extrapolating to the vertical axis at υ = 0 27All values of αn, α, n, E0′and k0 were represented in table 1

3.2.5. Square wave voltammetry (SWV)

Figure 8 shows the square wave voltammograms of the propsed electrode containing 0. 4 g Sephadex in a standard additions experiment .using BR buffer solution (pH =7) containing 0.03mM SDS, the solution was stirred for 30 s at 400 rpm at open circuit potential. The measurements were done at a peak height 50mV and pulse frequency 100 Hz.

3.3 Calibration, detection limit and reproducibility

A linear relationship between SLC anodic peak current of and its concentration was found in the range from 0.005mM to 1.0mM (R= 0.999) with a slop of 50.67µAmM -1and a limit of detection 0.001mM.The reproducibility (%RSD, n =3) of the peak current for 0.005mM sulfaclozine was 2.08% as shown in Table 2.

3.4 Determination of sulfaclozine in veterinary formulation

SLC was determined in Clozicocc® W.S.P. using standard addition method; the results obtained by the proposed method were statistically compared with those obtained by applying the reference method.5The calculated t- and F-values are found to be less than the theoretical ones, confirming accuracy and precision at 95% confidence level as shown in Table 3.

4 Conclusion

A novel simple square wave voltammetric method based on sephadex modified carbon paste electrode was developed for determination of sulfaclozine in veterinary formulations.  The electrochemical oxidation behavior of sulfaclozine was studied in a micellar medium by cyclic voltammetry. The electrochemical oxidation of SLC was found to be an irreversible process.An accumulation time of 30 s at open circuit potential was found to be useful for better sensitivity.Under the optimal conditions the SWV method was to be linear over a wide concentration range (5µM-1 mM SLC). The proposed method developed is simple and inexpensive for the determination of SLC in veterinary formulations with satisfactory accuracy and precision.

Figure 1: Chemical structure of Sulfaclozine

Figure 2: Cyclic voltammogram of 0.1 mM SLC in BR buffer of pH 7.0 at bare CPE.

Figure 3: Dependence of peak potential (a) and peak current (b) on

The pH of 0.1 mM SLC. The scan rate is 100 mV s-1

Figure 4: Cyclic voltammograms of 1.0 mM SLC in PH 7.0 using different modified and unmodified electrodes.

Figure 5: Cyclic voltammograms of SLC (1.0 mM) in PH 7.0 at different ratios of sephadex electrodes, at scan rate 100 mV s-1

Figure 6: Cyclic voltammograms recorded in 1.0 mM SLC containing different concentrations of SDS, the measurements were carried out using 40% sephadex carbon paste electrode in BR buffer pH 7.0.

Figure 7: (a) BR buffer solution of pH 7 containing 0.03 mM with different scan rates, (b) Dependence of the logarithm of peak current Ip/µA on logarithm of scan rate (υ /Vs−1). (c) Relationship between peak potential Ep/V and logarithm of scan rate log (υ/Vs−1). The scan rate was performed using an 40% SMCPE in 1.0 mM SLC containing 0.03 mM SDS in BR buffer of pH 7.0.

Figure 8: SWV voltammograms of SLC linearity (0.005mM to 1.0 mM) in PH (7.0) at 40% SMCPE &0.03 mM SDS

Table 1: The calculated values of αn, α, n, E0′ and k0 for the electro-oxidation of SLC by cyclic voltammetry (CV) at SMCPE.

| Parameters | SMCPE |
| --- | --- |
| αn | 0.8476 |
| α | 0.611 |
| n | 1.38 |
| E0′ | 0.85 |
| k0 | 3.1439 |

Table 2: Performance data of the proposed SWV method for determination of SLC

| SLC | Parameters |
| --- | --- |
| 0.005mM to 1.0 mM 50.67 5.86 0.9995 1.04 2.99 100.18 ± 0.01 2.08 2.12 | Linearity range ( mM ) Slope (µA. mM -1) Intercept (µA) Correlation coefficient (r) LOD ( µM ) LOQ( µM ) Accuracy (mean ± S.D.) Precision (RSD %) Interday Intraday |

Table 3: determination of SLC in pharmaceutical dosage form and statistical comparison of the proposed voltammetric and the published HPLC method 5

| Reference method5 | Standard addition technique |  |  |  | Pharmaceutical formulation |
| --- | --- | --- | --- | --- | --- |
|  | %Recoverya | Found(mg) | Added (mg) | Taken(mg) |  |
|  | 99.96 | 0.0149 | 0.015 | 0.015 | Clozicocc w.s.p batch No.564 |
|  | 98.87 | 1.43 | 1.45 | 0.120 |  |
|  | 100.31 | 2.01 | 2.00 | 0.120 |  |
| 99.33 | 99.59 |  |  |  | Mean |
| 1.51 | 1.01 |  |  |  | S.D |
| 5 |  |  |  |  | n |
| 2.28 | 1.02 |  |  |  | variance |
| ---- | 0.64 |  |  |  | Student's t-test (2.132)b |
| ---- | 2.24 |  |  |  | F-test (6.39)b |

a The results are the average of five separate determinations

b the tabulated t and F values, respectively, at P=0.05

5 References

- Chapman, H.D., Chapter 53 - Coccidiosis in Egg Laying Poultry, in Egg Innovations and Strategies for Improvements, P.Y. Hester, Editor. 2017, Academic Press: San Diego. p. 571-579.

- Chapman, H.D., Perspectives for the control of coccidiosis in poultry by chemotherapy and vaccinationin Proceedings of the IXth International Coccidiosis Conference. 2005. Foz de Iguassu, Parana, Brazil. p. 99-104.

- S.C. Sweetman, R.P.S., Martindale: The Complete Drug Reference. thirty-eights ed. Vol. 1. 2014, London: Pharmaceutical Press.

- Md. Harun-or-Rashid, et al., Scholars Journal of Agriculture and Veterinary Sciences, 2016. 3(4): p. 284-287.

- TANG Shu, C.J., GAO Jian-long,BAO En-dong, Nanjing Agricultural Univeristy, 2012: p. 105-109.

- Yu, H., et al., Journal of Chromatography B, 2011. 879(25): p. 2653-2662.

- Bousova, K. Senyuva, and H.  Mittendorf, Journal of Chromatography A, 2013. 1274: p. 19-27.

- Gorissen, B., et al., Analytical and Bioanalytical Chemistry, 2015. 407(15): p. 4447-4457.

- Goodspeed, et al., Journal - Association of Official Analytical Chemists, 1978. 61(5): p. 1050-1053.

- Su, H.X., et al., Fenxi Ceshi Xuebao, 2013. 32(2): p. 156-161.

- Ivan Švancara, et al.,Electroanalysis with Carbon Paste Electrodes. 2012: CRC Press Taylor & Francis Group.

- Kalcher, K., et al., Electroanalysis, 1995. 7(1): p. 5-22.

- Acuna, J.A., et al., Talanta, 1993. 40(11): p. 1637-42.

- R.Vittal, H. Gomathi, and K.J. Kim, Adv Colloid Interface Sci, 2006. 119(1): p. 55-68.

- Gel Filtration, Theory and Practice, Pharmacia Fine Chemicals. 1976, Uppsala:

Sweden.

- Radi.A, Analytical and bioanalytical chemistry, 2004. 378: p. 822-6.

- Radi.A, Fresenius' Journal of Analytical Chemistry, 1999. 364: p.590-594.

- ICH Q2A. validation of analytical methods. International Conference on Harmonization. 2003, IFPMA: Geneva.

- V. Momberg, et al., Analytica Chimica Acta, 1984. 159: p. 119-127.

- Sanghavi, B.J. and A.K. Srivastava, Electrochimica Acta, 2010. 55(28): p. 8638-8648.

- Jayaprakash, G.K., et al., Journal of Molecular Liquids, 2017. 240: p. 395-401.

- Manjunatha JG, et al., Int J Electrochem Sci, 2009. 4: 662-671.

- Shankar, S.S., B.E.K. Swamy, and B.N. Chandrashekar, Journal of Molecular Liquids, 2012. 168: p. 80-86.

- D.K. Gosser, Cyclic Voltammetry; Simulation and Analysis of Reaction Mechanisms.

New York (N.Y.) : VCH, 1993.

- E. Laviron, J. Electroanal. Chem. 1979. 101(1): p. 19-28

- A. J. Bard and L. R. Faulkner, “Electrochemical Methods Fundamentals and Applications,” 2nd Edition, Wiley, Hoboken, 2004.

- Wu, Y., X. Ji, and S. Hu, Bioelectrochemistry. 2004, 64,  p. 91-97