Synthesis of CuO by Electrospinning Method for Sensing of Hydrogen and Carbon Monoxide Gases
Abdollah Fallah Shojaei1*and Parisa Fallah Komsari1
1Department of Chemistry, Faculty of Science, University of Guilan, P. O. Box, 41335-1914, Rasht, Iran
*Corresponding author: E-mail: shoja47@yahoo.com, a.f.shojaie@guilan.ac.ir
ABSTRACT
The pure CuO nanofibers were synthesized via the electrospinning method successfully. The CuO nanoparticles were investigated for sensing hydrogen and carbon monoxide gases. Structural properties of the synthesized nanoparticles were studied using Fourier -transform infrared spectroscopy (FTIR), X-ray diffraction (XRD), Energy-dispersive X-ray spectroscopy (EDX), and particle morphology by scanning electron microscopy (SEM). SEM images confirmed string-like structures, nanofibers. The sensor based on the CuO nanoparticles exhibited excellent gas sensing performance at the low operating temperature of 175°C and the fast response and recovery characteristics at a low concentration. Moreover, good stability, prominent reproducibility, and excellent selectivity are also observed based on the nanoparticles. These results demonstrate the potential application of CuO nanoparticles for sensing H2 (10-200 ppm) and CO (400-700ppm).
Keywords: CuO, Carbon Monoxide, Electrospinning, Hydrogen, Nanoparticles
1. Introduction
  Metal oxide semiconductor (MOS) gas sensors have gained special focus driven by their diverse applications in air quality detection, inflammable gas inspection, environmental monitoring, healthcare, defense, security, and so on.1 Semiconducting metal oxides have been extensively used as sensing materials; resistance changes in oxide based semiconductor gas sensors are used to monitor reducing, toxic and inflammable gases, such as NH3, NO2, H2 and CO. 2 
In recent years, huge efforts have been made to develop nanostructured metal oxides with p-type semiconductivity. 3 Cupric oxide (CuO) as a p-type semiconducting oxide is known for being an excellent catalyst of Hydrogen gas.4 A variety of methods such as thermal oxidation,5 hydrothermal,6 and electrochemical deposition7 have been proposed and investigated for the fabrication of low dimensional CuO. Among these methods, electrospinning is one of the most attractive strategies for producing nanofibers of CuO. It has the merits of simplicity, high efficiency, low cost, and high reproducibility.8 In the other hand, many studies have reported on H2 and CO gases sensing. For example, in 1950, Wagner et al. reported the variation of electrical properties when ZnO is exposed to reducing gases. 9 After that, a series of research works about the sensing behavior of MOS to reducing gases were reported by Seiyama, et al. since the 1960s. 10 Up to present, many kinds of MOS were investigated as hydrogen and carbon monoxide sensing materials, including SnO2, TiO2, CuO, Al2O3, Bi2O3, Cr2O3, CuO, Fe2O3, NiO, TiO2, and ZnO which exhibit large variations in resistance after exposure to hydrogen and carbon monoxide gases. 11-36
In this paper, we reported a simple and facile approach to fabricate high quality CuO nanoparticles by electrospinning and their H2 and CO sensing characteristics are investigated. The study focuses on the design of CO and H2 sensors with minimum interference with other Volatile organic compounds. The high response, quick response, and recovery, and good selectivity are observed in our investigation, which indicates the potential application of CuO nanoparticles for the fabrication of high performance H2 and CO sensors.
2. Experimental 
2. 1. materials and methods 
Ethanol (>99%), N, N-dimethyl formamide (>99%, DMF), CuCl2⋅6H2O were used and purchased from Merck Company. Poly (vinyl pyrrolidone) (PVP, Mw = 1,300,000) and poly (vinyl alcohol) (PVA, Mw=31,000-50,000) were obtained from Sigma Aldrich and also distilled water was used. All chemicals were analytical grade and used as received without any further purification.
The Fourier transforms infrared spectroscopy (FT–IR) of the CuO nanoparticles were recorded on Thermo Nicolet Avatar spectrophotometer in the range of 500–4000 [image: image2.png]


using KBr pellets. The X-ray diffraction (XRD) technique (PHILIPS XRD diffractometer using Cu Kα (Kα = 1.540 Å and 2θ=10-90 radiation as X-ray source) was used to verify the structure of the CuO nanoparticles. Scanning electron microscopy (SEM) nanoparticles images were performed on an LEO 1450 VP (Germany) instrument and Energy Dispersive X-ray spectroscopy (EDX) was attained on a TESCAN S8000 microscope.
2. 2. Synthesis of CuO nanofibers
The electrospinning process was employed for CuO nanofibers synthesis. A mixed solution of polyvinyl alcohol (PVA, Mw: 31,000-50,000) and Copper (II) chloride (CuCl2. 6H2O) was used for electrospinning. In the typical experimental procedure, a PVA solution (10 wt %) was first prepared by dissolving PVA powder in distilled water and stirred 5 h at 60 °C. A 20 wt% CuCl2 solution and PVP powder was added to this solution. After stirring for 12 h, this mixed solution was introduced in 20 mL syringe equipped with a metal needle tip in a controlled electrospinning setup. The parameters for electrospinning were selected as; applied voltage: 20 kV, feeding rate of the solution: 0.2 mL/h, distance between glass substrate attached to the copper sheet, and the tip of the needle: 10 cm. Finally, the fibers were peeled off from the collector with tweezers and placed in a crucible. The conversion of copper dichloride to CuO and the removal of organic constituents PVP in the as-spun nanofibers were achieved by calcining at 400 °C for 2 h in air.
[image: image3.jpg]



[image: image4.jpg]CuO nanoparticles

Computer

Power Supply

’ =
g Mtimeter

Glass

Compartment

Sensor





Figure 1. Schematic image of (a) sensor kit, and (b) the gas-sensing analysis system.
2. 3. Gas identification system
Gas sensors were manufactured by screen-printing technique with planar glass substrates. Fig.1 (a) shows the schematic of the sensor. The as-prepared CuO nanoparticles were blended with deionized water in a weight ratio of 100: 25 to form a paste. Then the paste was subsequently screen-printed onto the planar glass substrate to form a sensing film with a thickness of about 60 𝜇m, the manufactured sensor was dried in air at room temperature to volatilize the water solvent and last aged in an aging test chamber for 24 h and after a Ni–Cr heating wire was inserted in the underneath to form a side-heated gas sensor. The efficiency of the gas sensors was determined in a gas testing chamber. Detail of the measurement setup can be observed in the schematic diagram as presented in Fig.1 (b). Electrical linkages to the gas sensors were achieved by physically connecting needle probes to the Pt electrodes of the sensors. A PC 510a (japan) multimeter was applied for measuring the changes of sensor resistance over time and a computer was used for logging data from the multimeter. At first, the responses of the sensors were measured in the presence of 200 ppm H2 gas and 700 ppm CO gas in the ambient air where the temperatures were modified in the range 25-175 °C in order to ascertain their optimum operating temperature. The sensors were pre-heated at different operating temperatures for about 45 min. When the resistances of all the sensors were stable, saturated target gas was injected into the test chamber (20 L in volume) by a microinjector through a rubber plug. After its resistance value reached a new constant value, the test chamber was opened to recover. The sensor resistance and sensitivity were collected and analyzed by the system. And the environmental temperature, relative humidity, and working temperature were recorded.
3. Results and Discussion
3. 1. Identification of sensor structure 

The X-ray diffraction method was used for the crystallization studies of the CuO nanoparticles. The X-ray diffraction pattern of the CuO fibers calcined at 400 °C for 2h (nanoparticles) is shown in Fig. 2. The CuO peaks appear at diffraction angles of 32/53°, 35/55°, 38/75°, 48/75°, 51/40°, 58/35°, 61/57°, 66/28°, 68/14°, 73/01°, 75/28°corresponding to reflection from (1 1 0), (0 0 2), (1 1 1), (-2 0 2), (0 2 0), (2 0 2), (-1 1 3), (-3 1 1), (2 2 0), (3 1 1), and (-2 2 2) planes, respectively. The strongest diffraction pattern viewed at 2θ =35/55 suggests that the CuO grows with a preferential orientation of (002) on the glass plate and illustrates the formation of the single-phase of monoclinic CuO. It was also viewed that the obtained XRD spectra are in very good agreement with reported 2θ values in JCPDS card no 48-1548.37

[image: image5.jpg](zzT)
(022)
(€1€)
€n-)
(z07)
[(11))
D
fzon) [(]1)] M
E 2 § £ &8 ¢

@ © -

Ayisudjuy dInjosqy

260

80.0

700

60.0

50.0

400

300

200

10.0




Figure 2. X-ray diffraction pattern of CuO nanofibers calcined at 400 °C for 2 h (nanoparticles).
Fig. 3(a) shows the SEM image of the CuO nanofibers, this precursor in the room temperature is with size ranging from 28 to 141 nm. After calcination for 2h at 400 °C, the size of the product, as shown in Fig. 3(b), it is thicker than that of the precursor. So that the nanofibers are crushed and the average diameter of the final product (CuO nanoparticles) is with size ranging from 37 to 315 nm.
[image: image6.jpg]



[image: image7.jpg]



Figure 3. SEM image of (a) CuO nanofibers (b) CuO nanofibers calcinated at 400 C for 2 h (nanoparticles).
    The FT-IR absorption peaks were performed to confirm the presence of characteristic vibrational peaks of CuO nanoparticles. The FT-IR spectrum of the CuO nanoparticles is presented in Fig. 4. The observed strong peaks at around 529, 585, and 672 [image: image9.png]


 are attributed to the stretching from Cu-O along (-202) direction. 38 The absorption peak around 1100 cm−1 may be attributed to –OH bending vibrations of Cu-OH.39 The transmittance peak that appeared at around 1377 cm−1 may be ascribed due to the presence of CO2, which is usually adsorbed from the air on the surface of sample materials during KBr pelletization.40 The absorption peak at1460.68 cm-1 may be due to, which usually appears in the spectrum when the FTIR sample are prepared and measured in the air. 41 The weak band at 1639.76 cm−1 may be attributed to the bending vibration modes of O–H groups of these adsorbed H2O molecules.41 Further, three small bands at 2862.62, 2928.94, and 2959.97 cm−1 may be ascribed to the asymmetric and symmetric vibrational modes of the O-H bond of the H2O molecules physisorbed on the surface of the CuO. 42 An intense broadband appeared in the 3400–3800 cm−1 region that was attributed to the O–H stretching vibration of surface hydroxyl groups of adsorbed water molecules, 43 which arises because nanocrystalline materials having a high surface-to-volume ratio absorbs high moisture.
[image: image10.png]1TysLe

zoz98T

Y68T6T

LE6S6T

LsTL01

— orszil

S
890981

— 9L6E91

— ovizre

00

[

6

8

6

6 96 s6
0/p@IUBY USR],

6

€6

3000 2500 2000 1500 1000
‘Wavenumber cm ™t

3500




Figure 4. FTIR spectra of CuO nanoparticles.
The EDX spectrum of the CuO nanoparticles is reported in Fig. 5. Only Cu and O signals have been detected, indicating that the nanoparticles were only built up of Cu and O. Weight percent of Cu was found to be 68.42 and that of O was 31.58. Thus the atomic ratio of Cu and O was 1:1.
[image: image11.png]Intensity (a.u.)

cuka
T

Energy (KeV)





Figure 5.  EDX spectra of CuO nanoparticles.
3. 2. Gas Sensing Performances
Despite the measurement simplicity required by MOX-based gas sensors, their detection mechanism is complex and not yet fully understood. Electrophysical and chemical properties, catalytic activity, thermodynamic stability, adsorption ability, the adsorption/desorption properties of the surface (highly dependent on materials used and environmental parameters) are just a few of the MOX parameters which are responsible for the sensing mechanism.44
The gas sensing properties of unloaded and loaded-MOS nanoparticles sensing films are characterized in terms of response, response time, and recovery time as a function of operating temperature, gas concentration, and metal loading. The resistance of the p-type metal oxide surface increases because generated electrons recombine with holes, decreasing the hole concentration. The response for p-type semiconducting oxide to reducing gas ([image: image13.png]SP



) becomes:
(1)                                                          Sp =[image: image15.png]



Where Rg and Ra are the electrical resistances of the sensors measured in the presence of reducing gas and pure dry air, respectively. The response time, Tres is defined as the time required reaching 90% of the steady response signal. The recovery times, Trec denotes the time needed to recover 90% of the original baseline resistance. Moreover, the selectivity defined as the response ratio of target gas to that of another gas is used to assess the relative performance of MOXs sensors towards different gases.45 In this paper the sensitivity (S) was defined as S = 𝑅g/𝑅a. 
[image: image16.png](a) In ambient air





[image: image17.png]®) Hy H,0

\ /‘ In ambient Hydrogen





Figure 6. The hydrogen sensing mechanism of resistance based MOS sensors
A possible sensing mechanism is described as follows to discover the gas sensing reaction process of CuO sensor against H2 and CO gases and illustrate the enhanced H2 and CO sensing properties of the nanoparticles.
 The resistance-based sensing mechanism of MOS is complicated and has been investigated by many researchers. 46 The commonly accepted mechanism is based on the variation of the surface electron depletion region due to the reaction between hydrogen and the chemisorbed oxygen on the surface. As described in Figure 6, under an air atmosphere the oxygen molecules can get adsorbed on the surface of the semiconductor and extracts electrons from the conduction band to form oxygen ions. That may lead to the formation of an electron depletion region near the surface, which can greatly increase the resistance due to the decrease of net carrier density. When the sensor is exposed to a hydrogen atmosphere, the hydrogen molecules will react with the adsorbed oxygen species. The redox reaction is exothermic and results in the fast desorption of produced H2O molecules from the surface. The released electrons will reduce the thickness of the depletion region, and decrease the resistance of the semiconductors. When the sensor is exposed to the air ambient again, the depletion region will be rebuilt by adsorbed oxygen species. The resistance will regain the initial level before hydrogen response.46
The adsorption of oxygen forms ionic species such as O2-, O- and O2-, which acquire electrons from the conduction band. The reaction kinetics may be explained by the following reactions: 47, 48
[image: image19.png]0,(g) & 0;(ad)



                                       (2)                                       
[image: image21.png]0;(ad) + e~ & 0; (ad)



                            (3)                              
[image: image23.png]0; + e~ & 20°



                                        (4)                                            
[image: image25.png]0 +e o 0%



                                              (5)                                      
H2 sensing mechanisms of the MOS can be explained from the following reaction paths: 49
[image: image27.png]H,(g)+ 0 (ad) - H,0+ e~



               (6)                        
[image: image29.png]H,(g) + 0° (ad) - H,0 + 2e~



             (7)        
[image: image30.png](a) In ambient air





[image: image31.jpg]In ambient

\ / Carbon monoxide





Figure 7. The Carbon Monoxide sensing mechanism of resistance based MOS sensors
As described in Figure 7, When the MOS gas sensors are exposed to air, oxygen molecules from the ambient air adsorb on the exposed MOS surface to form chemisorbed oxygen anions; ([image: image33.png]


,[image: image35.png]


 or [image: image37.png]


,depending on the operating temperature) by capturing electrons from the MOS conduction band. This results in the formation of a depletion layer, known as the space charge layer, on the surface of the MOS sensing material. This induces an increase in the resistance of the sensing materials. The reactions under various operating temperatures can be described as follows:
[image: image39.png]0,+e” =05



                         (8)
[image: image41.png]0, +2e” = 20~



                    (9)
[image: image43.png]0, +4e” = 20"



                (10)
When the MOS gas sensors are exposed to a CO environment at a moderate temperature, the adsorbed CO gas reacts with the chemisorbed oxygen anions on the MOS surface, resulting in CO2 molecules and the release of electrons back into the conduction band. This leads to an increase in the concentration of electrons. The resistivity of the MOS films decreases for detecting the CO gas. The reactions under various operating temperatures can be described as follows: 50 
[image: image45.png]2C0+0; = 2C0,+e”



                     (11)
CO + O-[image: image47.png]


CO2+ e-                                (12)
CO + O2-[image: image49.png]


CO2+ 2e-                            (13)
3. 3. Gas sensing characteristics
 In order to determine the optimum operating temperature, the responses of the CuO nanoparticles gas sensor to 200 ppm H2 and 700 ppm CO were measured continually at different operating temperatures. Fig. 8 shows the responses as a function of operating temperature from 25 to 300 °C.  For each gas, the response is measured to increase rapidly with increasing operating temperature and arrive at the maximum and then decreases with a further rise of the operating temperature. The optimum operating temperatures of the CuO nanoparticles are suggested at 175∘C for both H2 and CO gases with response values of 5.43 and 9.14, respectively.
[image: image50.jpg]Resistance (M€2)

200 ppm H,

150

Tempreture (°C)

200

250 300

350




[image: image51.jpg]=)
-

700 ppm CO

AN OWL TN N O

(BN dueIsIsaY

100 150 200 250 300 350
Tempreture (°C)

50




Figure 8. Responses of the CuO nanoparticles sensors to (a) 200 ppm of H2 and (b) 700 ppm of CO at different operating temperatures.
It is well known that response and recovery characteristics are important for evaluating the performances of semiconductor oxide sensors. The gas sensing transients of the CuO nanoparticles sensor upon exposure to H2 (200 ppm) and CO (700 ppm) are shown in Fig. 9. The gas responses and response speeds differed significantly according to the sensor temperature and type of gas. The times to reach 90% variation in resistance upon exposure to gas and air were defined as the 90% response time ([image: image53.png]


) and the 90% recovery time ([image: image55.png]


), respectively. Most of the [image: image57.png]


values were very short, indicating that both in the diffusion of analyte gas and the oxidation reaction between analyte gas and negatively charged chemisorbed oxygen occur very rapidly. Fast gas diffusion in the present study was attributed to the Porous structure of nanoparticles. fig.9 (a) shows the response time of the designed sensors for hydrogen and carbon monoxide gas was 39 s and 53 s at 175 °C, respectively. The recovery time for hydrogen and carbon monoxide gas was 94 s and 67 s 175 °C, respectively (fig.9 (b)).
[image: image58.jpg]<

LY

Apapisuag

250

200

150

100

50

Time(S)




[image: image59.jpg]67S

=)
=

Ao~ O WY MmN

Spanisuag

250

200

150

100

50

Time(S)




Figure 9.  A single-cycle response–recovery characteristic curves of CuO nanoparticles to (a) 200 ppm H2 and (b) 700ppm CO at 175° C.
To investigate the stability and repeatability of the CuO nanoparticles sensor, it was sequentially exposed to different levels of H2 gas (10, 25, 50,100, and 200 ppm) as shown in Figure 10 and shows the best response with more drastic resistance in concentration 200 ppm of H2. also Fig. 11 (a-d) presents three reversible cycles response curve in the concentration range of 400-700 ppm of the that confirmed the CuO nanoparticles had excellent stability and reversibility when alternately exposed to air and CO gas and it was found that CuO nanoparticles showed extremely excellent CO sensing performances at concentration 700 ppm with the highest sensor response and the highest selectivity compared to the other concentrations. As shown in Figures 10 and 11, the sensor response increases rapidly when exposed to a certain concentration of H2 and CO and decreases dramatically when exposed to air for recovery. Meanwhile, the gas response of the sensor always returns to its initial value during the continuous test period, implying a very satisfying reproducibility of the prepared sensor.
[image: image60.jpg]Sensitivity (Ry/R,)

18
1.6
14

12

0.8
0.6
0.4
0.2

(@)

10 ppm H,

1000

2000

3000
Time (S)

T

4000

5000

6000




[image: image61.jpg]Sensitivity (R/R,)
o S ~N w
ek N w s

o

(b)

25 ppm H,

0

100

200

300
Time (S)

400

500

600

700




[image: image62.jpg]50 ppm H,

©

4.5 4

4 4

n ooy N o
o ~ el

4y Ananisuag

St
=)

o

200 300 400 500 600 700

100

Time (S)




[image: image63.jpg]100 ppm H,

T T T
< L ~

a9 Sapsuag

200 300 400 500 600 700 800
Time (S)

100




[image: image64.jpg]200 ppm H,

< o0 ~

/) Ananisuag

-

200 300 400 500 600 700 800
Time (S)

100




Figure 10. Reproducibility of the CuO nanoparticles sensor on successive exposure (3 cycles) to (a) 10 ppm (b) 25 ppm (c) 50 ppm (d) 100 ppm (e) 200 ppm of H2 at 175°C.
[image: image65.jpg]Sensitivity (Rg/R,)

o o
NS

o o L7
o ®» B N B o

©

(a)

400 ppm CO

1000

2000

Time (S)

3000

4000 5000




[image: image66.jpg]500 ppm CO

—
°
=

wn

o
<

T My N
”m o~ —

“4/W) Amapisuag

10
5]

o

200 300 400 500 600 700 800

100

Time (S)




[image: image67.jpg]600 ppm CO

< ) ~

CuW) Ananisusg

300 400 500 600
Time ()

200

100




[image: image68.jpg]Sensitivity (R/R,)
O B N W & U1 O N O O

[
o

1@

700 ppm CO

100

200

300

400
Time (S)

500

600

T

700 800




Figure 11.  Reproducibility of the CuO nanoparticles sensor on successive exposure (3 cycles) to (a) 400 ppm (b) 500 ppm (c) 600 ppm (d) 700 ppm of CO at 175°C.
The concentration dependence of CuO nanoparticles was investigated in the concentration range of 10–200 ppm of H2 and the concentration range of 400-700 ppm CO the plots of the gas response against gas concentration are shown in Figure 12(a-b). As displayed in Figure12 (a), the gas response increases linearly with increasing the H2 concentration, which indicates the sensor is not saturated. Thus CuO nanoparticles sensor is favorable to detect H2 and CO with a low concentration in a log-log scale showed a linear relationship. 
[image: image69.jpg]Sensitivity (Ry/R,)

v

IS

w

[N}

-

o

50

100 150
Concentration (ppm)

200

250




[image: image70.jpg]Cco

Q @ ® ™~ VW T O N A O

/™) Ananisuds

200 300 400 500 600 700 800
Concentration (ppm)

100




Figure 12: Responses of the CuO nanoparticles sensor to different concentrations of (a) H2 and (b) CO at 175∘ C.
The gas sensing selectivity is another very important parameter to appraise the sensing ability of metal oxide semiconductor materials. Fig.13 describes the histogram of the gas response of the CuO nanoparticles sensor to 10 ppm of various gases, including n-hexane, methanol, ethanol, acetaldehyde, formaldehyde, and hydrogen at 175°C. This sensor shows an obvious hydrogen sensing response than other potential interface gases. Therefore, we believe that the CuO nanoparticles gas sensor has an excellent selectivity to10 ppm of H2.
[image: image71.jpg]Sensitivity




 
Figure 13. The selectivity of the CuO nanoparticles sensor on successive exposure to 10 ppm of various gases at 175°C.
4. CONCLUSION
In summary, Pure CuO nanofibers are synthesized via a simple electrospinning method and characterized by FT-IR, XRD, SEM, and EDX. The sensor exhibited the highest sensitivity (response: [image: image73.png]


/[image: image75.png]


= 5.43) to 200 ppm hydrogen and (response: [image: image77.png]


/[image: image79.png]


= 9.14) to 700 ppm Carbon monoxide at 175 °C. Short response and recovery times and excellent selectivity are indicated based on CuO nanoparticles that can be fabricated more quickly and at a lower cost. The simple method reported here can be used for further improving gas sensor performance such as enhancing the selectivity of the gas sensors to different test gases.
ACKNOWLEDGMENTS
This study was supported by the Science and Research Branch, University of Guilan. 
CONFLICT OF INTEREST
The authors declare that there are no conflicts of interest regarding the publication of this manuscript.
References
1. L. Li, C. Guo, S. Li, L. Wang, Q. Dong, W. Li, Sens. Actuators, B. 2010, 150, 806-810. ‏DOI:10.1016/j.snb.2010.07.022
2. C. C. Hsiao, L. S. Luo, Sensors, 2014, 14, 12219-12232. DOI:10.3390/s140712219 ‏
3. J. Jayaprakash, N. Srinivasan, P. Chandrasekaran, & E. K. Girija, Spectrochim. Acta, Part A. 2015, 136, 1803-1806. ‏DOI: 10.1016/j.saa.2014.10.087
4. N. D. Hoa, S. Y. An, N. D. Dung, N. V. Quy, D. Kim, Sens. Actuators, B. 2010, 146, 239-244. ‏ DOI:10.1016/j.snb.2010.02.045
5. D. T. T. Le, N. Van Duy, N. D. Hoa, N. Van Hieu, Phys. E (Amsterdam, Neth.). 2014, 58, 16-23. ‏DOI: 10.1016/j.physe.2013.11.013
6. M. A. Dar, Y. S. Kim, W. B. Kim, J. M. Sohn, H. S. Shin, Appl. Surf. Sci. 2008, 254, 7477-7481, ‏ DOI: 10.1016/j.apsusc.2008.06.004
7. L. Chen, S. Shet, H. Tang, H. Wang, T. Deutsch, Y. Yan, J. Turner, M. Al-Jassim, J. Mater. Chem. 2010, 20, 6962-6967. ‏ DOI: 10.1039/c0jm01228a
8. X. Shi, W. Zhou, D. Ma, Q. Ma, D. Bridges, Y. Ma, A. Hu, J. Nanomater. 2015, 16, 122. ‏DOI:10.1155/2015/140716
9. C. Wagner, J. Chem. Phys. 1950, 18, 69-71. ‏ DOI: 10.1063/1.1747459
10. T. Seiyama, A. Kato, K. Fujiishi, M. Nagatani, Anal. Chem. 1962, 34, 1502-1503.
11. M. Egashira, Y. Shimizu, Y. Takao, Sens. Actuators, B. 1990, 1, 108-112. ‏DOI: org/10.1016/0925-4005(90)80182-Y
12. B. Bott, T. A. Jones, B. Mann, Sens. Actuators. 1984, 5, 65-73. DOI:10.1016/0250-6874(84)87007-9
13. H. Nanto, T. Minami,  S. Takata,  J. Appl. Phys. 1986, 60, 482-484. DOI: 10.1063/1.337435
14. W. K. Choi, S. K. Song, J. S. Cho, Y. S. Yoon, D. Choi, H. J. Jung, S. K. Koh, Sens.  Actuators, B. 1997, 40, 21-27.  DOI: 10.1016/S0925-4005(97)80194-3
15. A. Galdikas, V. Jasutis, S. Kačiulis, G. Mattogno, A. Mironas, V. Olevano, D. Senulienė


 A. Šetkus, Sens. Actuators, B. 1997, 43, 140-146. DOI: 10.1016/S0925-4005(97)00206-2.
16. W. Y. Chung, G. Sakai, K. Shimano, N. Miura, D. D. Lee, N. Yamazoe, Sens. Actuators, B. 1998, 46,139-145.  DOI: 10.1016/S0925-4005(98)00100-2
17. I. Kosacki, H. U. Anderson, Sens. Actuators, B. 1998, 48, 263-269. DOI: 10.1016/S0925-4005(98)00055-0
18. Z. Tianshu, P. Hing, Y. Li, Z. Jiancheng, Sens. Actuators, B. 1999, 60, 208-215. DOI: 10.1016/S0925-4005(99)00272-5
19. J. H. Yu, G. M. Choi, Sens. Actuators, B. 2001, 75, 56-61. DOI: 10.1016/S0925-4005(00)00742-5
20. H. Gong, J. Q. Hu, J. H. Wang, C. H. Ong, F. R. Zhu, Sens. Actuators, B. 2006, 115, 247-251. DOI: 10.1016/j.snb.2005.09.008
21. A. Salehi, Sens. Actuators, B. 2003, 96, 88-93. DOI: 10.1016/S0925-4005(03)00490-8
22. P. Menini, F. Parret, M. Guerrero, K. Soulantica, L. Erades, A. Maisonnat, B. Chaudret, Sens. Actuators, B. 2004, 103, 111-114. DOI: 10.1016/j.snb.2004.04.103
23. C. H. Han, S. D. Han, I. Singh, T. Toupance, Sens. Actuators, B. 2005, 109, 264-269. DOI: 10.1016/j.snb.2004.12.115
24. T. Anukunprasert, C. Saiwan, E. Traversa, Sci. Technol. Adv. Mater. 2005, 6, 359-363. DOI: 10.1016/j.stam.2005.02.020
25. I. J. Kim, S. D. Han, C. H. Han, J. Gwak, H. D. Lee, J. S. Wang, Sensors, 2006, 6, 526-535. DOI: 10.3390/s6050526
26. L. Liu, C. Guo, S. Li, L. Wang, Q. Dong, W. Li, Sens. Actuators, B. 2010, 150, 806-810. DOI: 10.1016/j.snb.2010.07.022
27. J. K. Choi, I. S. Hwang, S. J. Kim, J. S. Park, S. S. Park, U. Jeong, Y. C. Kang, J. H. Lee, Sens.  Actuators, B. 2010,150, 191-199. DOI: 10.1016/j.snb.2010.07.013
28. J. H. Yoon, B. J. Kim, J. S. Kim, Mater. Chem. Phys. 2012, 133, 987-991.  10.1016/j.matchemphys.2012.02.002
29. S. Steinhauer, E. Brunet, T. Maier, G. C. Mutinati, A. Köck, O. Freudenberg, C. Gspan, W. Grogger, A. Neuhold, R. Resel, Sens. Actuators, B. 2013, 187, 50-57. DOI: 10.1016/j.snb.2012.09.034
30. E. Nikan, A. A. Khodadadi, Y. Mortazavi, Sens. Actuators, B. 2013, 184, 196-204. DOI: 10.1016/j.snb.2013.04.089
31. R. Ab Kadir, Z. Li, A. Z. Sadek, R. Abdul Rani, A. S. Zoolfakar, M. R. Field, J. Z. Ou, A. F. Chrimes, K. Kalantar-Zadeh, J. Phys. Chem. C. 2014, 118, 3129-3139. DOI: 10.1021/jp411552z
32. A. Harley-Trochimczyk, J. Chang, Q. Zhou, J. Dong, T. Pham, M. A. Worsley, R. Maboudian, A. Zettl, W. Mickelson, Sens. Actuators, B. 2015, 206, 399-406. DOI: 10.1016/j.snb.2014.09.057
33. C. C. Hsiao, L. S. Luo, Sensors, 2014, 14, 12219-12232. DOI: 10.3390/s140712219
34. O. Lupan, V. Postica, N. Ababii, M. Hoppe, V. Cretu, I. Tiginyanu, V. Sontea, Th. Pauporté,  B. Viana, R. Adelung, Microelectron. Engin. 2016, 164, 63-70. DOI: 10.1016/j.mee.2016.07.008
35. N. Xue, Q. Zhang, S. Zhang, P. Zong, F. Yang, Sensors, 2017, 17, 2351.DOI: 110.3390/s17102351
36. S. A. Mohd Chachuli, M. N. Hamidon, M. S. Mamat, M. Ertugrul, N. H. Abdullah, Sensors, 2018, 18, 2483. DOI: 10.3390/s18082483
37. Y. Zou, Y. Li, X. Lian, D. An, Research of Materials Science, 2014, 3, 44-51.
38. M. A. Das, S. H. Nam, Y. S. Kim, W. B. Kim, J. Solid State Electrochem.  2010, 14, 1719-1726. ‏DOI: 10.1007/s10008-010-1022-z
39. D. P. Dubal, G. S. Gund, C.D. Lokhande, R. Holze, Mater. Res. Bull. 2013, 48, 923-928. DOI: 10.1016/j.materresbull.2012.11.081
40. S. Sundar, G. Venkatachalam, S. J. Kwon, Nanomaterials, 2018, 8, 823. DOI: 10.3390/nano8100823
41. A. Nezamzadeh-Ejhieh, H. Zabihi-Mobarakeh, J. Ind. Eng. Chem. 2014, 20, 1421–1431. DOI: 10.1016/j.jiec.2013.07.027
42. P.K. Raul, S. Senapati, A. K. Sahoo, I. M. Umlong, R. R. Devi, A. J. Thakur, V. Veer, RSC Adv. 2014, 4, 40580-40587. DOI: 10.1039/C4RA04619F
43. M. Halder, M. M. Islam, Z. Ansari, S. Ahammed, K. Sen , S. K. M. Islam, ACS Sustain. Chem. Eng. 2017, 5, 648–657. DOI: 10.1021/acssuschemeng.6b02013
44. B.C. Serban, O. Buiu, C. Cobianu, M. Brezeanu, M. Bumbac, C.M. Nicolescu, Acta Chim. Slov. 2018, 65, 1014-1021. DOI: 10.17344/acsi.2018.4564.
45. T. Samerjai, N. Tamaekong, K. Wetchakun, V. Kruefu, C. Liewhiran, C. Siriwong, A. Wisitsoraat, S. Phanichphat, Sens. Actuators, B. 2012, 171, 43-61. DOI: 10.1016/j.snb.2012.05.052
46. H. Gu, Z. Wang, Y. Hu, Sensors, 2012, 12, 5517-5550. ‏DOI: 10.3390/s120505517
47. S. Phanichphant, C. Liewhiran, K. Wetchakun, A. Wisitsoraat, A. Tuantranont, Sensors, 2011, 11, 472–84. DOI: 10.3390/s110100472
48. L. F. Dong, Z. L. Cui, Z. K. Zhang, Nanostruct. Mater. 1997, 8, 815–823. DOI: 10.1016/S0965-9773(98)00005-1
49. C. Wongchoosuk, A. Wisitsoraat, D. Phokharatkul, A. Tuantranount, T. Kerdcharoen, Sensors, 2010, 10, 7705–15. DOI: 10.3390/s100807705
50. C. C. Hsiao, L. S. Luo, Sensors, 2014, 14, 12219-12232. DOI: 10.3390/s140712219
(a)





(b)





(a)





(b)





(a)





(b)





(a)





(b)










23



